## Evaluation of Active Passive Nature of Assistive Wearable Augmented Gait Suit for Enhanced Mobility

## Abstract

In this paper we discuss the design, control and active passive nature of the assistive wearable gait augment suit (AWGAS). AWGAS is designed to be soft, wearable, lightweight and assist walking gait by reducing muscle activation during walking. It augments walking by reducing the muscle activation of the posterior and anterior muscles of the lower limb. The suit uses pneumatic gel muscles (PGM), foot sensors for gait detection, and pneumatic valves to control the air pressure. The assistive force is provided using the motion in the loop feedforward control loop based on the foot sensors in shoes. PGMs are actuated with the help of pneumatic valves and portable air tanks. The elastic nature of the PGM allows AWGAS to assist walking in the absence of the air supply which makes AWGAS both active and passive walking assist suit. To evaluate the active and passive nature of the AWGAS, we experimented to measure surface EMG (sEMG) of the lower limb muscles. sEMG was recorded for unassisted walking, i.e., without the suit, passive assisted walking, i.e., wearing the suit with no air supply and active assisted walking, i.e. wearing the suit with air supply set at 80 kPa. The results show a reduction in the muscle activity for both passive and active assisted walking as compared to unassisted walking. The pilot trials of the AWGAS were conducted in collaboration with local farmers in the Hiroshima prefecture in Japan where feedback received is complementing the results obtained during the experiments.

## Introduction

Exoskeletons and exosuits augment human capability and reduce the chance of muscle fatigue, accidents or musculoskeletal disorders in the long term. These devices are used in settings such as factories where work involves lifting heavy objects, agriculture, walking long distances or sports. The objective of exoskeletons is to reduce human muscle efforts and augment the human ability to perform various tasks that could not be possible otherwise. Whereas, the objective of exosuits is to augment human motion and provide assistive force, thereby reducing muscle effort. With the growing elderly population, stressful work conditions, and eagerness to live a good quality of life, such devices will be critical.

Among various exoskeletons, HAL [1] makes walking easier and useful for the rehabilitation process. Wearable agri robot [2] designed to reduce physical strain on the body during farming activity. Walking assist device [3] reduces floor reaction force of the user and assist walking, stair climbing and squats motion. A passive plantarflexion assist exoskeleton [4] reduces the metabolic cost of walking. In stand-alone wearable power assist suits for caregivers, using pneumatic actuators [5], the suit provides extra strength and reduces the muscle effort required for the desired task up to 50\%. Exoskeleton to support knee joint motion [6] uses pneumatic artificial muscles to reduce the stress on the lower back and knees during the performance of loading and unloading tasks for manual workers. Wearable assistive suits for walking and lifting-up [7], with 13 DOF, and reduces muscle efforts significantly. A power assist device for standing up motion [8] for elderly or disabled to lead self-supporting lives. [9]–[12] are some of the wearable exosuits to assist walking and reduces the metabolic cost of walking using their unique methods.

These exoskeletons and exosuits are classified in various segments such as healthcare, disability support, industrial and augmented motion. They fulfill their respective purposes effectively using their unique methods, but the application outside the controlled environment is limited and not discussed especially for agriculture or outdoor environments. For augmenting human gait in an outdoor environment using exoskeleton or exosuits, wearable, portable, easy to use and light weight are some of the essential aspects to be considered, but together they are missing in the devices discussed earlier. These devices require large batteries or compressor or large air tanks for actuation of pneumatic artificial muscles (PAM), which makes device heavy and difficult to use. Most of these devices lack augmentation capability in the absence of the power sources such as batteries needed for actuator control or compressed air tanks in case of assistive suit using PAM. For the outdoor environment, it is essential that the device possess both active and passive nature to provide no disturbance or provide minimal assistance for human motion in the absence of power source. In our previous research [13], we developed a low-powered PAM, termed as pneumatic gel muscle (PGM). The PGM has an inner tube constructed with the styrene-based thermoplastic elastomer and outer protective mesh. This artificial muscle can be actuated with 50 kPa using small rubber pump and possess elastic nature to provide the minimal assistive force. Such actuation capacity is missing in the tradition PAM used for developing wearable assistive suits. In [13] K Ogawa et al. also devised a mechanism and developed an unplugged powered suit (UPS) to augment human walking.

Here, we discuss the enhanced design of the UPS using the motion in the loop feed-forward control algorithm for generating assistive force to improve the adverse effect on the antagonistic muscles discussed in [13]. In section 2 we will discuss the drawbacks of PGM and UPS, AWGAS design, biomechanics of the gait cycle, control algorithm and modeling of the assistive force generated by the PGM. In section 3, we discuss the evaluation and results of the AWGAS experiment. In section 4, we discuss preliminary trials in the rural area of the Hiroshima prefecture followed by discussion, conclusion to discuss the passive and active nature of the suit based on reduced muscle activity in the lower limb muscles.

## Background and Methodology

### Pneumatic Gel Muscle (PGM) and Unplugged Powered Suit (UPS)

In our previous research[13], we developed a low-pressure-driven pneumatic artificial muscle, PGM. Unlike McKibben PAM, the PGM is driven by low air pressure. K. Ogawa et al. reported that the PGM is elastic in nature and possesses the force-generating capacity from \SI{50}{\kilo\pascal} to \SI{300}{\kilo\pascal} of air pressure. Figure \ref{fig:pgmreal} shows the schematics and image of the actual product; the PGM used in this study has a natural length of \SI{300}{\milli\meter}, the maximum possible contraction is \SI{200}{\milli\meter}, the maximum elongation is \SI{500}{\milli\meter}, and the maximum possible force is \SI{50}{\newton}. K. Ogawa et al. in his study also designed and prototyped the UPS using PGM. The UPS requires no electricity but uses the walking motion to power the PGM with the help of the pump in the shoe. The primary evaluation of the UPS shows that it reduces the muscle efforts during walking by actuating PGM in the swing phase of the gait cycle; however, the effect is no significant for all the muscles of the lower limb owing to a small amount of air pressure generated by rubber pump, actuation delay due to pump and unstable walking due to pump in the shoe. For better performance, the use of external power sources such as an air tank and a pneumatic valve control can enhance the augmentation factor of the assistive suit.

<figure here>

To validate our assumption of use of air tank and pneumatic valves can eliminate the actuation delay, we conducted experiment to record and observe the delay due to rubber pump and air tank. We recorded air pressure and generated force with respect to time. Figure \ref{fig:activationdelayexperiment} shows the experimental setup; it consists of PGM with both ends fixed to the frame, air pressure sensor attached to the inlet valve of the PGM and loadcell force sensor attached to the other end. Table \ref{sensordetails} lists the specifications of the sensor. A rubber pump and a portable air tank was used for source of air pressure during the experiment. The actuation delay of PGM was observed for two configurations. First when PGM was attached at its rest length and second when it was stretched to \SI{370}{\milli\meter}. The result of the experiment are shown in figure \ref{fig:pump} and figure \ref{fig:tank} for pgm actuation using rubber pump and portable air tank. We observed that when the pump is used, the delay in the generated force is approximately \SI{260}{\milli\second} and \SI{130}{\milli\second} for the rest length of \SI{300}{\milli\meter} and the stretched length of \SI{370}{\milli\meter}, respectively. However, when the air tank is used as a source, no delay is observed between the supplied air pressure and generated force.

Based on this observation, we found that for UPS, generating the assistive force/torque precisely is difficult owing to the delay and the difficulty in attaching multiple pumps in the shoe, which could result in disturbing normal walking. Therefore, we developed a PGM actuation control based on the gait detection system that can benefit and elevate the performance of the walking assist suit. In the next section, we discuss the design and control of a newly developed walking assist suit.

### Design of Assistive Wearable Gait Augment Suit (AWGAS)

To overcome the challenge of the UPS, we developed a new walking assist suit called the AWGAS. AWGAS is developed based on the principle of the UPS, i.e., a lightweight, wearable, low-powered assistive suit. AWGAS uses pneumatic solenoid valves for the actuation control of the PGM to generate the assistive force and pneumatic regulator to control the supplied air pressure connected to portable air tank. The actuation control is designed based on the gait detection system developed using pressure sensors in the shoes. The new design enables the suit to support a walking pitch of more than two steps per second which was not possible in UPS\cite{11} and the level of augmentation can be controlled by setting the supply air pressure using regulator. K. Ogawa et al. mentioned that the maximum air pressure that the PGM can handle is 300 kPa, and PGM contraction increases with increase the supplied air pressure. Therefore, setting the maximum cutoff air pressure using air pressure regulator decides the augmentation factor of the suit during walking.

The AWGAS consists of a waist support, a knee support, PGMs attached along the quadriceps femoris, pressure sensors, pneumatic solenoid valves, controllers, portable air tank and air pressure regulator. Figure \ref{fig:aws} and figure \ref{fig:awgasillustration} show the overview and illustration of the applied assistive force during the swing phase of the gait cycle. Figure \ref{fig:aws} shows that the PGM is attached along the thigh muscles with the help of the waist support and knee support, while the backpack contains the controller circuit, the portable air tank connected to regulator, and a battery for the controller, and the pressure sensors in the shoe are connected to the controller. Figure \ref{fig:awgasillustration} illustrates the applied assistive force; the elastic nature of the PGM allows it to stretch during the stance and terminal stance phases and from the terminal stance controller that actuates the PGM for the swing phase of the gait cycle.

Table \ref{awgasconfig} lists all the components of the AWGAS with respective specification and weights. Based on this total weight of AWGAS is \SI{1.5}{\kilogram} approximately , the battery lasts for 3-4 days of runtime and the air tank can actuate 500 actuation at \SI{100}{\kilo\pascal} when tested in the lab.

<figure here>

The AWGAS uses the PGMs along the quadriceps femoris muscles connected across two joints: the pelvis and the knee. This configuration provides the assistive force during the swing phase of the gait cycle, and reduces the muscle effort of the quadriceps femoris during the swing phase of the gait cycle. The air pressure supply is controlled using the pneumatic solenoid valve. We also developed a gait detection system that identifies the swing phase of the individual limb. This system consists of two FSR-406 pressure sensors placed in each shoe. These are used together to detect the walking motion, standing posture, and gait cycle of the individual limb, and identify the phase of the gait cycle.

## AWGAS Evaluation Experiment with sEMG and Variable Assistive Force.

Gait cycle involves a complex sequence of the lower limb muscle activation together with joint moments at the hip, knee, and ankle. Augmenting the gait cycle involves reducing the muscle activation without affecting the walking speed or disturbing the motion. The AWGAS was developed to reduce muscle activation during the swing phase of the gait cycle by providing an assistive force using the PGM. The assistive force provided is based on the supplied air pressure. As discussed in the previous section, the AWGAS provides the assistive force for walking in the absence of air pressure. Based on this assumption, we evaluated the AWGAS for two conditions, one without air supply and the other with actuation control enabled and supply of air pressure. For the second case, the maximum possible air pressure was set to \SI{80}{\kilo\pascal}.

### Test Selection and Evaluation Criteria

To evaluate the changes in the lower limb muscle activation for To evaluate the changes in the lower limb muscle activation for the use of AWGAS, surface EMG (sEMG) of eight significant muscles involved in the gait cycle and superficially accessible to record the sEMG were evaluated. We recorded muscle activities of the biceps femoris (BF) of hamstring muscles, quadriceps femoris muscles i.e. rectus femoris (RF), vastus medialis (VM) and vastus lateralis (VL), the calf muscles such as soleus (SOL), lateral gastrocnemius (LG) and medial gastrocnemius (MG), and tibialis anterior (TA) muscle. These muscles are significant contributors to the gait cycle and commonly used for assessing the gait experiments.

For statistical analysis and comparison, multiple trials of the gait experiment with similar walking conditions were conducted, with each experiment measuring at least three full gait cycles excluding the beginning and the terminal gait cycle due to loading and unloading effect [19]. In our study, we conducted three trials of each experiment and recorded ten full gait cycles.

The experiment involves walking 15 m straight on a flat surface during the experiment sEMG, and the foot sensor data are recorded. Walking for 15 m provides ten full gait cycles, excluding the beginning and the terminal gait. For recording the sEMG, we used the P-EMG plus device with eight channels with the recording frequency of \SI{1}{\kilo\hertz} for all the channels. The sEMG data are synchronized with the help of the FSR sensor data, which is beneficial for segmenting the gait cycles for the analysis and identification of the assist period in the gait cycle.

Eight healthy adults seven male and one female (25.5 $\pm$ 4.8; mean $\pm$ s.d.) participated in the experiment. The details of the experiment were shared with all the participants before the experiment, and informed consent was acquired. During the experiment, the participants could rest anytime to avoid muscle fatigue during training and experiment session. Before starting the experiment, the maximum voluntary contraction (MVC) of each muscle was recorded. MVC is used for normalization of sEMG data for statistical analysis. To record the MVC, specific exercises or motions were performed by participants. They performed calf raises for recording the MVC for the calf muscles, i.e., SOL, LG, and MG, dorsiflexion for TA and squats for BF, and thigh contraction for VM, VL, and RF.

After measuring MVC, participants were trained to maintain walking speed and get used to walking with the assistive suit. First, each subject walked 15 m repeatedly without an assistive suit and trained to maintain the walking speed. During this training, participants got familiar with the walking speed and surroundings. After this, participants wear AWGAS and practice walking to get familiar to walk with AWGAS. During which participants trained for assistive walking by wearing AWGAS with and without assistive air pressure provided. For each experiment the participants trained for approximately 20 mins.

* First each subject walked 15 m without assistive suit and trained to maintain the walking speed. During this training, subjects got familiar with the walking speed and experiment surrounding.
* After this, subjects put on the AWGAS and practice walking to get familiar to walk with AWGAS. This experiment was conducted for wearing AWGAS only and with assistive air pressure provided.
* The time required for training varies for each participant.
* For each experiment the participants trained for approximately 20 mins.
* Between each training session participants were asked to rest for 10 to 15 mins
* AWGAS is soft walking assist suit and does not restricts the normal walking when not powered.

Figure \ref{fig:experiment} shows a participant wearing the experiment setup. The participant wears the AWGAS, with the sEMG electrodes attached to respective muscles, shoes with the FSRs, a backpack, a P-EMG device for recording the sEMG, a laptop for connecting and operating the P-EMG application, a battery for powering the P-EMG device and AWGAS controller, a controller and actuation circuit for the AWGAS, and a portable air tank. Together with all the equipment, the weight of the backpack is 7 kg. The laptop in the backpack was remotely operated for recording the sEMG appropriately.

## Preliminary Trials and Assessment by Elderly Farmer

The AWGAS is developed to augment walking and to be used by people of all age groups. To study the feasibility and usefulness of the AWGAS, we conducted the preliminary trials, where we shared the details of the AWGAS and its purpose with people in rural areas and requested to test the AWGAS and provide feedback. For successful participation, informed consent was obtained from all participants and pertinent information such as AWGAS specification and documentation, purpose of the trial, procedure, benefit and risks was translated in Japanese language and shared 4 months before the trials.

During the trials, participants walked with and without AWGAS on paddy field while performing various daily rural tasks. They walked on a flat surface, up the slope, down the slope, and on uneven surfaces with and without the AWGAS. Figure \ref{fig:preliminarystudy} shows the participant testing the AWGAS. After the trial, participants said they feel less effort required in the lower limb when using assistive suit as compared to walking without the suit. They mentioned about feeling of assistive force while walking but are not sure how it reduces the efforts. They also mentioned that the backpack is lightweight and does not feel load on the back, but most farmers need to visit paddy fields using small trucks. In such cases, wearing and removing the suit due to backpack is time consuming. In our evaluation experiment we observed AWGAS reduces lower limb’s muscle efforts while walking. The feedback received during trials complemented the evaluated results. This proves that we can use AWGAS not in controlled environment but also in the uncontrolled environment. During trial we identified two areas of improvements. First making compact design will improve usability for their daily lifestyle. Second, the current system uses air tank which lasts for 500 actuations as described in section 2.2. During the trial we used two air tanks so that it lasts longer. We feel, a better management of air tanks or strategy to replace the tanks is needed to support large number of actuations.

~~The feedback received from the participant was that the AWGAS did not interfere with the various tasks attempted and walking with the AWGAS was easier as compared to walking without the AWGAS. The result of our laboratory experiment showed a significant reduction in the muscle activity. In the preliminary field study, we received feedback that indicated the ease of walking with the AWGAS in an uncontrolled environment. In the future, we plan to conduct more of such experiments to gather more feedback and improve the device.~~

During the trials, participants walked with and with AWGAS on paddy field while performing various daily rural tasks.

After the trial, participants said they feel less effort required in the lower limb when using assistive suit as compared to walking without the suit. They also mentioned about feeling of assistive force while walking but are not sure how it reduces the efforts. They also mentioned that the weight of the backpack is lightweight, but most farmers need to visit paddy fields using small trucks. In such case wearing and removing the suit is time consuming. In our evaluation experiment we observed AWGAS reduces lower limb’s muscle efforts while walking. The feedback received during trials complemented the evaluated results. This proves that we can use AWGAS not in controlled environment but also in the uncontrolled environment. During trial we identified two areas of improvements. First making compact design will improve usability for their daily lifestyle. Second, the current system uses air tank which lasts for 500 actuations as described in section 2.2. During the trial we used two air tanks so that it lasts longer. We feel, a better management of air tanks or strategy to replace the tanks is needed to support large number of actuations.

Points to be added

* Effects of backpack
  + The weight of the AWGAS is 1.2 kg as discussed in Section 2. The backpack used during the trial contains the controller circuit and portable compressed air tank. Based on the feedback received from the trials the backpack seems very lightweight and did not disturb normal walking. But many farmers have paddy fields at various location and travel between them using small truck. The current version of AWGAS does not support users have to remove the backpack which will disconnect the controller and air tank from artificial muscles.
  + We will try to improve this shortfall in the design in our future work.
* Battery life hour
  + The purpose of the battery is to power the controller and pneumatic solenoid valves. The specification of the battery are as follows
    - Capacity:10400 mAh / 38.48 wh
    - Output:5V / 3A
    - Input:5V / 2A
    - Weight:8.5oz / 240g
    - Size:3.8 x 3.1 x 0.9in / 97 x 80 x 22mm
  + Specification of solenoid valves
    - Voltage: +5V
    - Current: 100 mA
    - Power Consumption: 1 W
  + Total runtime for the battery
    - Runtime = (10\*Ampere Hours) / Load in watts
    - Runtime = (10\*10)/1 = 100 hours approximately
* Air tank life hour
  + Air tanks are used for actuation of PGM actuator
  + Air tank capacity Gas quantity : 30-40 liters
  + At 100 kPa 3 actuation per 1 gm of the of air tank i.e. 74\*3 = 222 actuation (gait cycle) [this is sue’s data]
  + At 100 kPa one full tank goes can be used for 1 km

Above points are added in AWGAS design section to describe the missing information about the suit. A new table is added to describe these in section 2.2

In Section 4 i.e. preliminary trials we describe the feedback from the users, similarity with the lab results, challenges and future improvements in design to overcome the shortfall in current design and assist capabilities.

## References

[1] K. Suzuki, G. Mito, H. Kawamoto, Y. Hasegawa, and Y. Sankai, “Intention-based walking support for paraplegia patients with Robot Suit HAL,” *Adv. Robot.*, vol. 21, no. 12, pp. 1441–1469, 2007.

[2] S. Toyama and G. Yamamoto, “Development of wearable-agri-robot - Mechanism for agricultural work,” in *2009 IEEE/RSJ International Conference on Intelligent Robots and Systems, IROS 2009*, 2009, pp. 5801–5806.

[3] Y. Ikeuchi, J. Ashihara, Y. Hiki, H. Kudoh, and T. Noda, “Walking assist device with bodyweight support system,” in *2009 IEEE/RSJ International Conference on Intelligent Robots and Systems, IROS 2009*, 2009, pp. 4073–4079.

[4] P. Malcolm, W. Derave, S. Galle, and D. De Clercq, “A Simple Exoskeleton That Assists Plantarflexion Can Reduce the Metabolic Cost of Human Walking,” *PLoS One*, vol. 8, no. 2, 2013.

[5] M. Ishii, K. Yamamoto, and K. Hyodo, “Stand-Alone Wearable Power Assist Suit - Development and Availability,” *J. Robot. Mechatronics*, vol. 17, no. 5, pp. 17–18, 2005.

[6] Y. Naruoka, N. Hiramitsu, and Y. Mitsuya, “A study of power-assist technology to reduce body burden during loading and unloading operations by support of knee joint motion,” *J. Robot. Mechatronics*, vol. 28, no. 6, pp. 949–957, 2016.

[7] K. Sano, E. Yagi, and M. Sato, “Development of a Wearable Assist Suit for Walking and Lifting-Up Motion Using Electric Motors,” *J. Robot. Mechatronics*, vol. 25, no. 6, pp. 923–930, 2013.

[8] T. Noritsugu, D. Sasaki, M. Kameda, A. Fukunaga, and M. Takaiwa, “Wearable power assist device for standing up motion using pneumatic rubber artificial muscles,” *J. Robot. Mechatronics*, vol. 19, no. 6, pp. 619–628, 2007.

[9] B. T. Quinlivan *et al.*, “Assistance magnitude versus metabolic cost reductions for a tethered multiarticular soft exosuit,” vol. 4416, no. January, 2017.

[10] A. T. Asbeck, K. Schmidt, and C. J. Walsh, “Soft exosuit for hip assistance,” *Rob. Auton. Syst.*, vol. 73, pp. 102–110, 2015.

[11] K. Schmidt *et al.*, “The myosuit: Bi-articular anti-gravity exosuit that reduces hip extensor activity in sitting transfers,” *Front. Neurorobot.*, vol. 11, no. OCT, pp. 1–16, 2017.

[12] A. T. Asbeck, S. M. M. De Rossi, K. G. Holt, and C. J. Walsh, “A biologically inspired soft exosuit for walking assistance,” *Int. J. Rob. Res.*, vol. 34, no. 6, pp. 744–762, 2015.

[13] K. Ogawa, C. Thakur, T. Ikeda, T. Tsuji, and Y. Kurita, “Development of a pneumatic artificial muscle driven by low pressure and its application to the unplugged powered suit,” *Adv. Robot.*, vol. 31, no. 21, pp. 1135–1143, 2017.